## README
Illumination power at the sample plane — bench worksheet
Mirrors kb/calibrations/illumination-power.yaml. Fill the YELLOW cells only.
What this unblocks, and what it does not
power_at_sample_mw is {} for every line of every source today, so
optics.path.excitation_rate_per_s returns None and lens 5 is BLOCKED by design.
UNBLOCKED by this measurement: G21 (light-driving) · G22 (total dose)
STILL BLOCKED afterwards: G10 (photobleaching)
G10 needs bleach_photons, which is per dye and empty for every dye. A power meter
cannot supply it. Afterwards, say "G10 only" — not "lens 5 works now".
Two traps, both checked on the Checks tab before any mW is written
1. The spot may be larger than the sensor.
A head reads only what lands on its element. If the illuminated field at 4x
overfills it, every 4x row is low by an unknown factor and nothing downstream
can tell.
2. A photodiode head is wavelength-calibrated.
Reading 561 nm with the meter still set to 488 gives a plausible number that is
wrong by the responsivity ratio. A thermal head does not care — record which.
Why 13 rows per line and not 144
For one monochromatic line, transmission of mirror x multiband x IR is exactly the
product of the three, so the 2x2x2 factorial is redundant. The cross-check rows test
that the product actually holds. If it does not, that is the more important result
and the reduction is void — the Derived tab computes the comparison for you.
Level linearity does NOT factor and must be swept: it decides linear_in_level in
data/light_sources.yaml, whose schema stores one mW at 100% assuming it holds.
The half a power meter does not give
optics.path.excitation_rate_per_s needs BOTH power_mw_at_sample AND
illuminated_area_um2; the gates consume the ratio, W/cm2. The magnification effect
lives mostly in the area term — the field diaphragm's image scales as 1/M, so area
goes as 1/M^2 and 4x to 20x is a factor of ~25, against a few percent of
transmission. Measuring power at three objectives without the Area tab measures the
small term three times.
The objective that is not on this sheet
The standing condition is 100x Oil, and 4x/10x/20x do not reach it.
Transmission does not extrapolate across objectives. If 100x is not measured,
power_at_sample_mw stays empty for the objective the real work uses, and lens 5
stays BLOCKED there. Measuring it means immersion oil between the front element and
the meter window; reading it dry gives a number, and a wrong one.
Decide: measured this session / deferred / out of scope. Any is fine; silence is not.
Where these numbers go afterwards
mW at 100%, per line per objective -> data/light_sources.yaml lines.<line>.power_at_sample_mw
whether the % sweep was linear -> data/light_sources.yaml lines.<line>.linear_in_level
element in/out ratios -> a cross-check on lens 1's spectral model, not a new field
illuminated area -> the illuminated_area_um2 argument optics.path already takes
Do not fill a cell from a datasheet, a percentage, or another objective.
An empty cell is a valid result. A wrong one is not.

## Meta
Provenance — a measurement without this is not re-usable
docs/02 §1: date and measurer are mandatory for anything in kb/calibrations/.
Field | Value | What it means
date |  | YYYY-MM-DD
measured_by |  | who was at the bench
meter_model
head_type |  | thermal ¦ photodiode
head_wavelength_set |  | nm the meter was set to — note if changed per line
head_aperture_mm |  | sensor diameter (trap 1)
head_position |  | e.g. "on the stage, at the sample plane"
mm_config |  | which .cfg was loaded
room_light |  | on ¦ off ¦ shrouded
100x_oil_decision |  | measured ¦ deferred ¦ out of scope
notes

## Checks
Pre-checks — these gate every row on the Power tab
1. Does the illuminated spot fit inside the sensor?
If FALSE for an objective, every mW at that objective is a LOWER BOUND, not a measurement.
Objective | Fits? (TRUE/FALSE)
4x
10x
20x
2. Is 0% the same as the shutter closed?
Any difference is leakage, and belongs in the record rather than in the noise.
Reading | mW
at 0% level
shutter closed
difference (leakage)
3. Which state is the instrument's NORMAL running config?
Every ratio on the Derived tab is taken against it.
mirror |  | in ¦ out
multiband |  | in ¦ out
ir |  | in ¦ out
objective |  | e.g. 20x

## Power
Power at the sample plane — 13 rows per line
Fill column J (mW) and, where useful, K. Everything else is the condition, already decided.
Blue row = the reference every ratio is taken against. Grey row = worked example, do not fill.
Block | Purpose | Source | Line (nm) | Objective | Mirror | Multiband | IR | Level % | mW | Notes
(example) | REFERENCE | Aura | 488 | 20x | in | in | in | 100 | 1.234 | head centred, room dark
Line 1 | linearity |  |  | 20x | in | in | in | 0
Line 1 | linearity |  |  | 20x | in | in | in | 5
Line 1 | linearity |  |  | 20x | in | in | in | 10
Line 1 | linearity |  |  | 20x | in | in | in | 30
Line 1 | linearity |  |  | 20x | in | in | in | 50
Line 1 | REFERENCE |  |  | 20x | in | in | in | 100
Line 1 | objective |  |  | 4x | in | in | in | 100
Line 1 | objective |  |  | 10x | in | in | in | 100
Line 1 | element |  |  | 20x | out | in | in | 100
Line 1 | element |  |  | 20x | in | out | in | 100
Line 1 | element |  |  | 20x | in | in | out | 100
Line 1 | cross-check |  |  | 20x | out | out | in | 100
Line 1 | cross-check |  |  | 20x | out | out | out | 100
Line 2 | linearity |  |  | 20x | in | in | in | 0
Line 2 | linearity |  |  | 20x | in | in | in | 5
Line 2 | linearity |  |  | 20x | in | in | in | 10
Line 2 | linearity |  |  | 20x | in | in | in | 30
Line 2 | linearity |  |  | 20x | in | in | in | 50
Line 2 | REFERENCE |  |  | 20x | in | in | in | 100
Line 2 | objective |  |  | 4x | in | in | in | 100
Line 2 | objective |  |  | 10x | in | in | in | 100
Line 2 | element |  |  | 20x | out | in | in | 100
Line 2 | element |  |  | 20x | in | out | in | 100
Line 2 | element |  |  | 20x | in | in | out | 100
Line 2 | cross-check |  |  | 20x | out | out | in | 100
Line 2 | cross-check |  |  | 20x | out | out | out | 100
Line 3 | linearity |  |  | 20x | in | in | in | 0
Line 3 | linearity |  |  | 20x | in | in | in | 5
Line 3 | linearity |  |  | 20x | in | in | in | 10
Line 3 | linearity |  |  | 20x | in | in | in | 30
Line 3 | linearity |  |  | 20x | in | in | in | 50
Line 3 | REFERENCE |  |  | 20x | in | in | in | 100
Line 3 | objective |  |  | 4x | in | in | in | 100
Line 3 | objective |  |  | 10x | in | in | in | 100
Line 3 | element |  |  | 20x | out | in | in | 100
Line 3 | element |  |  | 20x | in | out | in | 100
Line 3 | element |  |  | 20x | in | in | out | 100
Line 3 | cross-check |  |  | 20x | out | out | in | 100
Line 3 | cross-check |  |  | 20x | out | out | out | 100
Line 4 | linearity |  |  | 20x | in | in | in | 0
Line 4 | linearity |  |  | 20x | in | in | in | 5
Line 4 | linearity |  |  | 20x | in | in | in | 10
Line 4 | linearity |  |  | 20x | in | in | in | 30
Line 4 | linearity |  |  | 20x | in | in | in | 50
Line 4 | REFERENCE |  |  | 20x | in | in | in | 100
Line 4 | objective |  |  | 4x | in | in | in | 100
Line 4 | objective |  |  | 10x | in | in | in | 100
Line 4 | element |  |  | 20x | out | in | in | 100
Line 4 | element |  |  | 20x | in | out | in | 100
Line 4 | element |  |  | 20x | in | in | out | 100
Line 4 | cross-check |  |  | 20x | out | out | in | 100
Line 4 | cross-check |  |  | 20x | out | out | out | 100

## Area
Illuminated area — the half a power meter does not give
Image a uniform fluorescent slide or dye solution; measure the illuminated disc's diameter in pixels.
In data/pixel_size.yaml only 20x is evidence: measured. 4x and 10x are nominal, so their area inherits nominal unless the pixel size is measured in the same session (python -m calibration.cli intermediate-mag <cfg>).
Objective | Diameter (px) | um/px | px evidence | Shape | Area (um^2) | Area (cm^2)
4x
10x
20x
nominal ¦ measured — write it, the tier travels with the number
Shape: disc ¦ vignetted ¦ field-filling. A vignetted field is not a disc of that diameter.

## Derived
Computed at the bench — nothing here is filled by hand
Blank until the Power tab has values. Watch the cross-check while you measure.
Block 1
Linearity |  |  |  |  | linear_in_level is TRUE only if these deviations stay small
Level % | Measured mW | If linear | Deviation
0
5
10
30
50
100
Element transmission |  |  |  |  | T = reference / (that element removed). T<1 means it attenuates.
Element | mW with it out | T
mirror
multiband
IR
Cross-check: does the product hold? |  |  |  |  | If agreement departs from 1.00, the 13-row reduction is void.
Case | Predicted mW | Measured mW | Agreement
mirror + multiband out
all three out
Per objective, and the number the gates actually eat
Objective | mW @100% | vs 20x | Area (cm^2) | Irradiance (W/cm^2)
4x
10x
20x
Block 2
Linearity |  |  |  |  | linear_in_level is TRUE only if these deviations stay small
Level % | Measured mW | If linear | Deviation
0
5
10
30
50
100
Element transmission |  |  |  |  | T = reference / (that element removed). T<1 means it attenuates.
Element | mW with it out | T
mirror
multiband
IR
Cross-check: does the product hold? |  |  |  |  | If agreement departs from 1.00, the 13-row reduction is void.
Case | Predicted mW | Measured mW | Agreement
mirror + multiband out
all three out
Per objective, and the number the gates actually eat
Objective | mW @100% | vs 20x | Area (cm^2) | Irradiance (W/cm^2)
4x
10x
20x
Block 3
Linearity |  |  |  |  | linear_in_level is TRUE only if these deviations stay small
Level % | Measured mW | If linear | Deviation
0
5
10
30
50
100
Element transmission |  |  |  |  | T = reference / (that element removed). T<1 means it attenuates.
Element | mW with it out | T
mirror
multiband
IR
Cross-check: does the product hold? |  |  |  |  | If agreement departs from 1.00, the 13-row reduction is void.
Case | Predicted mW | Measured mW | Agreement
mirror + multiband out
all three out
Per objective, and the number the gates actually eat
Objective | mW @100% | vs 20x | Area (cm^2) | Irradiance (W/cm^2)
4x
10x
20x
Block 4
Linearity |  |  |  |  | linear_in_level is TRUE only if these deviations stay small
Level % | Measured mW | If linear | Deviation
0
5
10
30
50
100
Element transmission |  |  |  |  | T = reference / (that element removed). T<1 means it attenuates.
Element | mW with it out | T
mirror
multiband
IR
Cross-check: does the product hold? |  |  |  |  | If agreement departs from 1.00, the 13-row reduction is void.
Case | Predicted mW | Measured mW | Agreement
mirror + multiband out
all three out
Per objective, and the number the gates actually eat
Objective | mW @100% | vs 20x | Area (cm^2) | Irradiance (W/cm^2)
4x
10x
20x